## Tree diagram & P-Diagram
Medical
Adapter
Mechanical
功能：提供單體之結構保護、散熱、絕緣及電力傳輸。
Case
功能：提供單體外部結構強度保護、符合規格之尺寸與外觀。
Heatsink
功能：提供晶體散熱、單體內部零件電磁屏蔽及結構功能。
Insulator
功能：提供單體高低壓側絕緣功能及導熱(Inulator Pad)功能。
Electrical
功能：提供PSU主要電氣功能 (Made by EE....)
Cable
功能：負責單體 輸入及輸出之功能。
Cover(Plastic) 功能：提供單體外部結構強度保護、符合規格之尺寸與外觀之組合零件。
Chassis((Plastic) 功能：提供單體外部結構強度保護、符合規格之尺寸與外觀之組合零件。
Primary Side HS 功能：提供Bridge Diode & MosFET散熱功能。
Secondary Side HS 功能：提供二次側MosFET散熱功能。
Shieding HS 功能：提供單體散熱及電磁屏蔽功能。
Insulator Pad 功能：提供晶體導熱及絕緣功能。
Insulator Sheet 功能：提供單體絕緣功能。
Input Cable 功能：將市電傳導至單體輸入側之功能。
Insulator Sheet 功能：將單體轉換完成之電力傳送至負載端。

## DFMEA
Design Failure Mode and Effects Analysis (DFMEA)
PLANNING AND PREPARATION (STEP 1 規劃與準備)
| Company Name: | ABC 公司 | Subject: | Power Supply(Adapter)
| Engineering Location: | 台北市 Taiwan | DFMEA Start Date: | Nov.2 , 2023 | DFMEA ID Number: | DFMEA-1
| Customer Name: | EFG 客戶 | DFMEA Revision Date: | Nov.20 , 2024 | Design Responsibility: | RD
| Model Year¦ Platform: | Medical Adapter 90W | Cross-Functional Team: | EE, ME, RD | Confidentiality Level: | Medium
| STRUCTURE ANALYSIS (STEP 2 結構分析) |  |  | FUNCTION ANALYSIS (STEP 3 功能分析) |  |  | FAILURE ANALYSIS (STEP 4 失效分析) |  |  |  | RISK ANALYSIS (STEP 5 風險分析) |  |  |  |  | OPTIMIZATION (STEP 6 優化) | [12 empty] | CONTINUOUS IMPROVEMENT 持續改善
Issue # 議題號碼 | 1. Next Higher Level 上一階 | 2. Focus Element 標的物 | 3. Next Lower Level 下一階 or Characteristic Type (Geometry, Material, Surface Finish, Coating, etc.) | 1. Next Higher Level Function and Requirement 上一階功能需求 | 2. Focus Element Function and Requirement 標的物功能需求 | 3. Next Lower Level Function and Requirement of Characteristic 下一階功能需求 | 1. Failure Effects (FE) to the Next Higher Level Element and/or End User 上一階失效效應 | 2. Failure Mode (FM) of the Focus Element 標的物失效模式 | 3. Failure Cause (FC) of the Next Lower Element or Characteristic 下一階失效原因 | Original Severity (S) of FE 嚴重度 | Current Prevention Control (PC) of FC 目前預防管制 | Occurrence (O) of FC 發生度 | Current Detection Control (DC) of FC or FM 目前偵測管制 | Detection (D) of FC/FM 難檢度 | DFMEA AP 優先行動 | DFMEA Prevention Action 未來預防行動 | DFMEA Detection Action 未來偵測行動 | Responsible Person 負責單位 | Target Completion Date 預計完成日 | Status: (Open, Completed, Discarded) 狀態 | Action Taken with Pointer to Evidence 佐證 | Completion Date 實際完成日 | Severity (S) | Occurrence (O) | Detection (D) | DFMEA AP | Remarks 備註 | Issue # | History/Change Authorization 歷史資料/授權
1 | Adapter(Electrical) | Input & EMI Stage | Input Fuse | Provide 24Vdc/90W(Max.) stable output voltage & current | Handle enough input power, EMI suppress performance & safety requirements & Rectifier AC input voltage to positive voltage. | Protect input mains voltage from adapter power stage short circuit. | Adapter output shutdown | Fuse blows with normal operation condition. | Fuse Current rating or I2t rating is too low | 9 | Current rating stress shal l<80%, I2t stress shall <70% | 2 | Calculation & function check | 2 | L | None
2 | Adapter(Electrical) | Input & EMI Stage | Bleeder Resistor | Provide 24Vdc/90W(Max.) stable output voltage & current | Handle enough input power, Suppress EMI noise & meet safety requirements & Rectifier AC input voltage to positive voltage. | Discharge X-cap. Residual voltage after remove input voltage | User might get electronic shock if they touch AC power cord terminal after adapter remove AC input. | X-cap. discharge time over safety requirements. | Resistance is too high. | 9 | Time constant of bleed resistor & all X-cap. shall <1sec. | 2 | Calculation & function check | 2 | L | None
3 | Adapter(Electrical) | Input & EMI Stage | X-Cap. | Provide 24Vdc/90W(Max.) stable output voltage & current | Handle enough input power, Suppress EMI noise & meet safety requirements & Rectifier AC input voltage to positive voltage. | Suppress EMI noise | EMI fail with regulatory requirements. | Conducted EMI test fail. | Capacitance of X-Cap. is too low. | 9 | The capacitance follow standard design/best practices. | 4 | PSU EMI Test | 2 | H | None | Final product of test: measuring adapter conducted EMI with capacitance close to lower limit value. | EE | 2024-08-31 | Open |  |  | 9 | 2 | 4 | L
4 | Adapter(Electrical) | Input & EMI Stage | Bridge Diode | Provide 24Vdc/90W(Max.) stable output voltage & current | Handle enough input power, Suppress EMI noise & meet safety requirements & Rectifier AC input voltage to positive voltage. | Rectifier AC input voltage to positive voltage. | Adapter output shutdown | Bridge diode become short circuit. | Ifsm rating is not enough | 8 | Ifsm rating meet circuit analysis & de-rating analysis requirement. | 4 | Component electrical stress test. | 2 | M | None
5 | Adapter(Electrical) | PFC stage | Boost Inductor | Provide 24Vdc/90W(Max.) stable output voltage & current | Increase the power factor as close as 1. With power factor correction, the input peak current can be reduced, thus reducing components rating requirement/cost and increasing system efficiency. | The inductor could store/release energy to boost the output voltage as desired level. | The adapter efficiency become lower, thus increase the adapter temperature and decrease the adapter lifetime. | PFC stage couldn't correct the power factor because of the shorted inductor. | The boost inductor is too hot and damaged as a shorted wire (with no inductance), duo to too high wiring resistance (result in too high power loss) of the inductor wire. | 9 | The inductor winding diameter shall meet circuit analysis requests, and the thermal simulation results shall fulfill de-rating requirements. | 2 | Apply design review and thermal simulation. | 1 | L | None
6 | Adapter(Electrical) | PFC stage | MOSFET | Provide 24Vdc/90W(Max.) stable output voltage & current | Increase the power factor as close as 1. With power factor correction, the input peak current can be reduced, thus reducing components rating requirement/cost and increasing system efficiency. | By turning on/off the MOSFET, the energy could be stored/released in the boost inductor, thus an desired output voltage could be established. | The adapter been shutdown, no output. | PFC stage output voltage range out of spec, also PFC couldn't support enough holdup time for the adapter output. | The MOSFET is too hot and damaged as a shorted wire, duo to too high Cross capacitance (result in too high power loss) of the MOSFET wire. | 9 | The MOSFET Coss rating and other critical parameters of the MOSFET shall meet circuit analysis request, and the thermal simulation results shall fulfill de-rating requirements. | 3 | Apply design review and thermal simulation. | 1 | L | None
7 | Adapter(Electrical) | PFC stage | (Blocking) Diode | Provide 24Vdc/90W(Max.) stable output voltage & current | Increase the power factor as close as 1. With power factor correction, the input peak current can be reduced, thus reducing components rating requirement/cost and increasing system efficiency. | Working as an one way valve to prevent the output energy from flowing back to input power source, thus an desired output voltage could be established. | The adapter been shutdown, no output. | PFC stage couldn't work because of the broken diode. | The diode is damaged as an open(broken) component, duo to too high inrush current passing through the diode. | 9 | The current rating and other critical parameters of the diode shall meet circuit analysis requirements. | 4 | Apply design review and component stress test. | 2 | H | None | EVT stage product of test: measuring the inrush current under worst case conditions. | EE | 2024/11/31 | Open |  |  | 9 | 2 | 2 | L
8 | Adapter(Electrical) | PFC stage | Bulk Capacitor | Provide 24Vdc/90W(Max.) stable output voltage & current | Increase the power factor as close as 1. With power factor correction, the input peak current can be reduced, thus reducing components rating requirement/cost and increasing system efficiency. | The bulk capacitor could store/release energy to keep the output voltage stable, and provide holdup-time by demand. | The adapter lost its essential function (holdup time & PFC output voltage out of spec). | PFC stage output voltage range out of spec, also PFC couldn't support enough holdup time for the adapter output. | The bulk capacitor is too hot that the capacitance was decreased abnormally, duo to too high ripple current of the capacitor. | 8 | The ripple current rating and other critical parameters of the bulk capacitor shall meet circuit analysis request, and the thermal simulation results shall fulfill de-rating requirements. | 3 | Apply design review and thermal simulation. | 6 | M | None | EVT stage product of test: measuring the ripple current and temperature under worst case conditions. | EE | 2024/11/31 | Open |  |  | 8 | 3 | 2 | L
9 | Adapter(Electrical) | DC/DC Stage | Primary MOSFET | Provide 24Vdc/90W(Max.) stable output voltage & current | Convert high voltage (144V~400V) to rated voltage (24V) | MOSFET turn on and turn off to transport energy | Adapter output shutdown | MOSFET become short-circuit. | The spike voltage sis over MOSFET Drain-Source max. rated voltage | 8 | Perform circuit analysis and calculation, according to the calculated value, the derating values of MOSFET are selected. | 4 | Circuit Calculation and Component electrical stress test. | 4 | M | None | DQ Test | EE | 2024-11-02 | Open |  |  | 8 | 2 | 2 | L
10 | Adapter(Electrical) | DC/DC Stage | Secondary MOSFET | Provide 24Vdc/90W(Max.) stable output voltage & current | Convert high voltage (144V~400V) to rated voltage (24V) | Rectify and regulate output voltage and current with primary MOSFET | Adapter output shutdown | MOSFET become short-circuit. | The spike voltage sis over MOSFET Drain-Source max. rated voltage | 8 | Perform circuit analysis and calculation, according to the calculated value, the derating values of MOSFET are selected. | 4 | Circuit Calculation and Component electrical stress test. | 3 | M | None
11 | Adapter(Electrical) | DC/DC Stage | Transformer | Provide 24Vdc/90W(Max.) stable output voltage & current | Convert high voltage (144V~400V) to rated voltage (24V) | Energy storage and transfer | EMI fail | Transformer coupling change to poor | Transformer core crack or transformer foot pin fracture | 9 | Transformed follow standard design & best practice. | 2 | Vibration test and structural simulation | 2 | L | None
12 | Adapter(Electrical) | DC/DC Stage | Output Capacitor | Provide 24Vdc/90W(Max.) stable output voltage & current | Convert high voltage (144V~400V) to rated voltage (24V) | Keep stable output voltage | Unstable output voltage | Capacitor characteristic deterioration or open circuit | Long time under high temperature causes electrolyte evaporation | 7 | Keep away from heating component as much as possible | 2 | Thermal simulation and test | 3 | L | None
13 | Adapter(Electrical) | Feedback & Control Stage | TEA1753T VIN_SENSE PIN Resistor | Provide 24Vdc/90W(Max.) stable output voltage & current | Detection input voltage and current to provide PFC/Flyback MOSFET Driver correct control signal. Enter protection mode when the output occurs overvoltage or overload condition. | AC input voltage divide | Adapter can not turn-on. | Resistance in high impedance, control IC can not start up. | Voltage rating is not enough. | 8 | Voltage rating shall meet circuit analysis and de-rating analysis requirements. | 4 | Design review & calculation | 4 | M | None | Change to high voltage CHIP Resistors. | EE | 2024-12-31 | Completed |  |  | 8 | 1 | 4 | L
14 | Adapter(Electrical) | Feedback & Control Stage | TEA1753T VCC PIN Capacitor | Provide 24Vdc/90W(Max.) stable output voltage & current | Detection input voltage and current to provide PFC/Flyback MOSFET Driver correct control signal. Enter protection mode when the output occurs overvoltage or overload condition. | Provide stable voltage, refrain VCC voltage drop | Adapter can not turn-on. | Capacitor failure, control IC can not start up. | Voltage rating is not enough. | 8 | Voltage rating shall meet circuit analysis and de-rating analysis requirements. | 2 | Circuit calculation & component stress test. | 2 | L | None
15 | Adapter(Electrical) | Feedback & Control Stage | TEA1753T PFCSENSE PIN Resistor | Provide 24Vdc/90W(Max.) stable output voltage & current | Detection input voltage and current to provide PFC/Flyback MOSFET Driver correct control signal. Enter protection mode when the output occurs overvoltage or overload condition. | PFCSENSE PIN Filter function | Power factor poor. | Resistance become high impedance, no control signal to PFC. | Power rating is not enough. | 7 | Power rating shall meet circuit analysis and de-rating analysis requirements. | 4 | Circuit calculation & component stress test. | 4 | M | None
16 | Adapter(Mechanical) | Case | Cover & Chassis | 提供單體之結構保護、符合規格之尺寸與外觀、散熱、絕緣及電力傳輸。 | 提供單體外部結構強度保護、符合規格之尺寸與外觀。 | 提供單體外部結構強度保護、符合規格之尺寸與外觀之組合零件 | 單體外殼破裂、內部零件損毀或功能異常 | 外殼破裂造成內部零件受外力損毀，造成單體功能異常 | Cover & Chassis 強度設計不足。(材質選用 & 厚度設計) | 8 | 利用結構模擬&結構測試確認強度需求，改善機殼結構強度(變更材質或增加厚度) | 4 | 球擊測試 &落下測試:樣品需經過球擊測試 &落下測試其結果不得違反安全之情形特別是外殼破裂距離不足與絕緣破損 | 2 | M | 設計初期利用結構模擬依照產品需求選用機殼材質與定義機殼厚度 | Worse case 之測試評估/依設計需求落實機殼零件厚度與材質之進料檢驗 | Test Engineer | 2024-11-02 | Open |  |  | 8 | 3 | 2 | L
17 | Adapter(Mechanical) | Case | Cover & Chassis | 提供單體之結構保護、符合規格之尺寸與外觀、散熱、絕緣及電力傳輸。 | 提供單體外部結構強度保護、符合規格之尺寸與外觀。 | 提供單體外部結構強度保護、符合規格之尺寸與外觀之組合零件 | 單體外觀尺寸不符規格 | 超音波熔接後溢膠，導致單體尺寸與外觀不良 | Cover & Chassis熔接線與超音波熔接參數配合設計不當/熔接操作不當 | 8 | 增加熔接參數設定確認流程/操作者加強訓練 | 2 | 加強目檢與尺寸檢驗 | 2 | L | Cover & Chassis熔接線與超音波熔接參數配合設計最佳化 | 熔接流程、外觀與尺寸檢驗標準化 | Test Engineer | 2024-11-02 | Open |  |  | 8 | 2 | 2 | L
18 | Adapter(Mechanical) | Heatsink | Shielding HS | 提供單體之結構保護、符合規格之尺寸與外觀、散熱、絕緣及電力傳輸。 | 提供晶體散熱、單體內部零件電磁屏蔽及結構功能。 | 提供單體散熱及電磁屏蔽功能。 | 單體內部主變壓器燒毀 | 散熱片導熱不良 | Shielding 與變壓器點膠間隙過大(熱阻高)，導致變壓器熱無法導出而積熱燒毀 | 8 | 測量分析間隙尺寸，變更組合零件尺寸或增加點膠量改善間隙過大問題 | 4 | 量測變壓器溫度確保散熱正常 | 3 | M | 利用Thermal Simulation與組裝零件尺寸公差分析計算出最佳化之點膠間隙 (取熱與結構之平衡) | 於研發各階段確實量測變壓器溫度之變化 | ME/EE/ Thermal Engineer | 2024-11-02 | Open |  |  | 8 | 3 | 3 | L
19 | Adapter(Mechanical) | Insulator | Insulator Sheet | 提供單體之結構保護、符合規格之尺寸與外觀、散熱、絕緣及電力傳輸。 | 提供單體高低壓側絕緣功能及導熱(Insulator Pad)功能 | 提供晶體導熱及絕緣功能。 | 單體內部零件功能異常 | arcing | 絕緣片破裂或絕緣片設計尺寸不足 | 7 | 確認PCB零件出腳長度，縮短出腳長度避免零件腳刺破絕緣片/依據跳火位置增加絕緣片尺寸 | 2 | HI-POT 測試確認 | 2 | L | 利用零件組裝尺寸公差分析PCB零件出腳空間限制，定義產品零件出腳長度/依照設計準則設計絕緣片尺寸 | 依照機種需求可在B/I測試後在加測HI-POT 2 測試確認 | ME/EE/ Thermal Engineer | 2024-11-02 | Open |  |  | 7 | 2 | 1 | L

## DFMEA Severity
Product General Evaluation Criteria Severity (S) 產品通用評估標準 嚴重性 S
Potential Failure Effects rated according to the criteria below. 潛在故障影響根據以下標準進行評級 |  |  | Blank until filled in by user 留空，直到使用者填寫為止
S | Effect 影響 | Severity criteria 嚴重性標準 | Corporate or Product Line Examples
10 | Very High 非常高 | Affects safe operation of the vehicle and/or other vehicles, the health of operator or passenger(s) or road users or pedestrians. 影響車輛和/或其他車輛的安全操作、操作人員或乘客或道路消費者或行人的健康。
9 |  | Noncompliance with regulations. 不符合規定。
8 | High 高 | Loss of primary vehicle function necessary for normal driving during expected service life. 在預期的使用壽命內喪失主要正常駕駛所需的基本車輛功能。
7 |  | Degradation of primary vehicle function necessary for normal driving during expected service life. 在預期的使用壽命內，正常駕駛所必需的主要車輛功能退化。
6 | Moderate 中等 | Loss of secondary vehicle function. 輔助車輛功能喪失
5 |  | Degradation of secondary vehicle function. 輔助車輛功能退化。
4 |  | Very objectionable appearance, sound, vibration, harshness, or haptics. 非常令人反感的外觀、聲音、振動、粗糙度或觸覺。
3 | Low 低 | Moderately objectionable appearance, sound, vibration, harshness, or haptics. 中度令人反感的外觀、聲音、振動、粗糙度或觸覺。
2 |  | Slightly objectionable appearance, sound, vibration, harshness, or haptics. 輕微令人反感的外觀、聲音、振動、粗糙度或觸覺。
1 | Ver Low 非常低 | No discernible effect. 沒有明顯的效果。

## DFMEA Occurrence
Occurrence Potential (O) for the Product 產品通用評估標準 發生性 O
Potential Failure Causes rated according to the criteria below. Consider Product Experience and Prevention Controls when determining the best Occurrence estimate (Qualitative rating) 潛在故障原因根據以下標準進行評級。 在確定最佳發生估計值（定性評級）時，請考慮產品體驗和預防控制 |  |  | Blank until filled in by user 留空，直到使用者填寫為止
O | Prediction of Failure Cause Occurring 故障原因的預測 | Occurrence criteria - DFMEA 發生標準 - DFMEA | Corporate or Product Line Examples
10 | Extremely high 極高 | First application of new technology anywhere without operating experience and / or under uncontrolled operating conditions. No product verification and/or validation experience. Standards do not exist and best practices have not yet been determined. Prevention controls not able to predict field performance or do not exist. 在沒有操作經驗和/或不受控制的操作條件下首次應用新技術。沒有產品驗證和/或確認經驗。 標準不存在，最佳做法尚未確定。預防控制措施無法預測現場性能或不存在。
9 | Very high 非常高 | First use of design with technical innovations or materials within the company. New application or change in duty cycle / operating conditions. No product verification and/or validation experience. Prevention controls not targeted to identify performance to specific requirements. 首先在公司內部使用具有技術創新或材料的設計。新應用或佔空比/操作條件的變化。沒有產品驗證和/或確認經驗。 預防控制不以識別特定要求的性能為目標。
8 |  | First use of design with technical innovations or materials on a new application. New application or change in duty cycle / operating conditions. No product verification and/or validation experience. Few existing standards and best practices, not directly applicable for this design. Prevention controls not a reliable indicator of field performance. 首次在新的應用中使用具有技術創新或材料的設計。新應用或佔空比/操作條件的變化。沒有產品驗證和/或確認經驗。 現有的標準和最佳實踐很少，不能直接適用於此設計。預防控制不是現場性能的可靠指標。
7 | High 高 | New design based on similar technology and materials. New application or change in duty cycle / operating conditions. No product verification and/or validation experience. Standards, best practices, and design rules apply to the baseline design, but not the innovations. Prevention controls provide limited indication of performance. 基於類似技術和材料的新設計。新應用或佔空比/操作條件的變化。沒有產品驗證和/或確認經驗。 標準、最佳實踐和設計規則適用於基線設計，但不適用於創新。預防控制提供的性能指示有限。
6 |  | Similar to previous designs, using existing technology and materials. Similar application, with changes in duty cycle or operating conditions. Previous testing or field experience. Standards and design rules exist but are insufficient to ensure that the failure cause will not occur. Prevention controls provide some ability to prevent a failure cause. 與以前的設計類似，使用現有的技術和材料。類似的應用，但佔空比或操作條件會發生變化。以前的測試或現場經驗。 標準和設計規則存在，但不足以保證不會發生故障原因。預防控制提供了一些功能來防止故障原因。
5 | Moderate 中 | Detail changes to previous design, using proven technology and materials. Similar application, duty cycle or operating conditions. Previous testing or field experience, or new design with some test experience related to the failure. Design addresses lessons learned from previous designs. Best Practices re-evaluated for this design but have not yet been proven. Prevention controls capable of finding deficiencies in the product related to the failure cause and provide some indication of performance. 使用成熟的技術和材料對以前的設計進行詳細更改。類似的應用、佔空比或操作條件。以前的測試或現場經驗，或具有一些與故障相關的測試經驗的新設計。 設計解決了從以前的設計中吸取的經驗教訓。已針對此設計重新評估了最佳實踐，但尚未得到證實。預防控制能夠發現產品中與故障原因相關的缺陷，並提供一些性能指示。
4 |  | Almost identical design with short-term field exposure. Similar application, with minor change in duty cycle or operating conditions. Previous testing or field experience. Predecessor design and changes for new design conform to best practices, standards, and specifications. Prevention controls capable of finding deficiencies in the product related to the failure cause and indicate likely design conformance. 幾乎相同的設計，具有短期的現場曝光。類似的應用，但佔空比或操作條件略有變化。以前的測試或現場經驗。 前代設計和新設計的更改符合最佳實踐、標準和規範。預防控制能夠發現產品中與故障原因相關的缺陷，並指示可能的設計一致性。
3 | Low 低 | Detail changes to known design (same application, with minor change in duty cycle or operating conditions) and testing or field experience under comparable operating conditions, or new design with successfully completed test procedure. Design expected to conform to Standards and Best Practices, considering Lessons Learned from previous designs. Prevention controls capable of finding deficiencies in the product related to the failure cause and predict conformance of production design. 對已知設計（相同應用，佔空比或操作條件略有變化）和可比操作條件下的測試或現場經驗的詳細更改，或成功完成測試程式的新設計。設計應符合標準和最佳實踐，同時考慮從以前的設計中吸取的經驗教訓。預防控制能夠發現產品中與故障原因相關的缺陷，並預測生產設計的一致性。
2 | Very low 非常低 | Almost identical mature design with long term field exposure. Same application, with comparable duty cycle and operating conditions. Testing or field experience under comparable operating conditions. Design expected to conform to Standards and Best Practices, considering Lessons Learned from previous designs, with significant margin of confidence. Prevention controls capable of finding deficiencies in the product related to the failure cause and indicate confidence in design conformance. 幾乎相同的成熟設計，具有長期的現場暴露。相同的應用，具有可比的佔空比和工作條件。在可比操作條件下的測試或現場經驗。考慮到從先前設計中吸取的經驗教訓，設計應符合標準和最佳實踐，並具有很大的信心度。預防控制能夠發現與故障原因相關的產品缺陷，並表明對設計一致性的信心。
1 | Extremely low 極低 | Failure eliminated through preventive control and failure cause is not possible by design. 通過預防性控制消除故障和故障原因在設計上是不可能的。
Product Experience: History of product usage within the company (Novelty of design, application or use case). Results of already completed detection controls provide experience within the design. 產品使用經驗：公司內部的產品使用歷史（設計、應用或用例的新穎性）。已完成的檢測控制結果為設計提供了經驗。
Prevention Controls: Use of Best Practices for product design, Design Rules, Company Standards, Lessons Learned, Industry Standards, Material Specifications, Government Regulations and effectiveness of prevention oriented analytical tools including Computer Aided Engineering, Math Modeling, Simulation Studies, Tolerance Stacks and Design Safety Margins 預防控制：在產品設計中使用最佳實踐、設計規則、公司標準、吸取的經驗教訓、行業標準、材料規格、政府法規以及以預防為導向的分析工具的有效性，包括計算機輔助工程、數學建模、模擬研究、公差堆棧和設計安全裕度
Note: O 10, 9, 8, 7 can drop based on product validation activities. 注意：O 10、9、8、7 可能會根據產品驗證活動而下降。

## DFMEA Detection
Detection Potential (D) for the Validation of the Product Design 產品通用評估標準 難檢性 D
Detection Controls rated according to Detection Method Maturity and Opportunity for Detection. 檢測控制根據檢測方法成熟度和檢測機會進行評級。 |  |  |  | Blank until filled in by user 留空，直到使用者填寫為止
D | Ability to Detect 偵測能力 | Detection Method Maturity 偵測成熟度 | Opportunity for Detection 偵測機會 | Corporate or Product Line Example
10 | Very Low 非常低 | Test procedure yet to be developed. 未定義測試步驟 | Test method not defined 未定義測試方法
9 |  | Test method not designed specifically to detect failure mode or cause. 測試方法不是專門為檢測故障模式或原因而設計的。 | Pass-Fail, Test-to-Fail, Degradation Testing 通過/不過測試, 失效測試,老化測試
8 | Low 低 | New test method; not proven 新的測試方法;未經證實 | Pass-Fail, Test-to-Fail, Degradation Testing 通過/不過測試, 失效測試,老化測試
7 |  |  | Pass-Fail Testing 通過/不過測試
6 | Moderate 中 | Proven test method for verification of functionality or validation of performance, quality, reliability and durability; planned timing is later in the product development cycle such that test failures may result in production delays for re-design and/or re-tooling. 經過驗證的測試方法，用於驗證功能或驗證性能、品質、可靠性和耐用性;計劃的時間安排在產品開發周期的後期，因此測試失敗可能會導致重新設計和/或重新加工工具的生產延遲。 | Test-to-Failure 失效測試
5 |  |  | Degradation Testing 老化測試
4 | High 高 | Proven test method for verification of functionality or validation of performance, quality, reliability and durability; planned timing is sufficient to modify production tools before release for production. 經過驗證的測試方法，用於驗證功能或驗證性能、品質、可靠性和耐用性;計劃的時間足以在發佈到生產環境之前修改生產工具。 | Pass-Fail Testing 通過/不過測試
3 |  |  | Test-to Failure 失效測試
2 |  |  | Degradation Testing 老化測試
1 | Very high 非常高 | Prior testing confirmed that failure mode or cause cannot occur, or detection methods proven to always detect the failure mode or failure cause. 先前的測試證實了故障模式或原因不會發生，或者檢測方法被證明始終能夠檢測故障模式或故障原因。

## DFMEA AP
Action Priority (AP) for DFMEA and PFMEA
Action Priority is based on combinations of Severity, Occurrence, and Detection ratings in order to prioritize actions for risk reduction. 行動優先順序是基於嚴重度S、發生度O和難檢度D評級的組合 為了確定降低風險的優先行動的優先次序。 | [6 empty] | Blank until filled in by user 留空，直到使用者填寫為止
Effect 效應 | S | Prediction of Failure Cause Occurring 預測 失敗原因 發生 | O | Ability to Detect 偵測能力 | D | ACTION PRIORITY (AP) 行動優先 | Comments
Product or Plant Effect Very high 產品或 工廠 影響 非常高 | 9-10 | Very high | 8-10 | Low-Very low | 7-10 | H
|  |  |  | Moderate | 5-6 | H
|  |  |  | High | 2-4 | H
|  |  |  | Very high | 1 | H
|  | High | 6-7 | Low-Very low | 7-10 | H
|  |  |  | Moderate | 5-6 | H
|  |  |  | High | 2-4 | H
|  |  |  | Very high | 1 | H
|  | Moderate | 4-5 | Low-Very low | 7-10 | H
|  |  |  | Moderate | 5-6 | H
|  |  |  | High | 2-4 | H
|  |  |  | Very high | 1 | M
|  | Low | 2-3 | Low-Very low | 7-10 | H
|  |  |  | Moderate | 5-6 | M
|  |  |  | High | 2-4 | L
|  |  |  | Very high | 1 | L
|  | Very low | 1 | Very high - Very low | 1-10 | L
Product or Plant Effect High 產品或 工廠 影響 高 | 7-8 | Very high | 8-10 | Low-Very low | 7-10 | H
|  |  |  | Moderate | 5-6 | H
|  |  |  | High | 2-4 | H
|  |  |  | Very high | 1 | H
|  | High | 6-7 | Low-Very low | 7-10 | H
|  |  |  | Moderate | 5-6 | H
|  |  |  | High | 2-4 | H
|  |  |  | Very high | 1 | M
|  | Moderate | 4-5 | Low-Very low | 7-10 | H
|  |  |  | Moderate | 5-6 | M
|  |  |  | High | 2-4 | M
|  |  |  | Very high | 1 | M
|  | Low | 2-3 | Low-Very low | 7-10 | M
|  |  |  | Moderate | 5-6 | M
|  |  |  | High | 2-4 | L
|  |  |  | Very high | 1 | L
|  | Very low | 1 | Very high - Very low | 1-10 | L
Product or Plant Effect Moderate 產品或 工廠 影響 中等 | 4-6 | Very high | 8-10 | Low-Very low | 7-10 | H
|  |  |  | Moderate | 5-6 | H
|  |  |  | High | 2-4 | M
|  |  |  | Very high | 1 | M
|  | High | 6-7 | Low-Very low | 7-10 | M
|  |  |  | Moderate | 5-6 | M
|  |  |  | High | 2-4 | M
|  |  |  | Very high | 1 | L
|  | Moderate | 4-5 | Low-Very low | 7-10 | M
|  |  |  | Moderate | 5-6 | L
|  |  |  | High | 2-4 | L
|  |  |  | Very high | 1 | L
|  | Low | 2-3 | Low-Very low | 7-10 | L
|  |  |  | Moderate | 5-6 | L
|  |  |  | High | 2-4 | L
|  |  |  | Very high | 1 | L
|  | Very low | 1 | Very high - Very low | 1-10 | L
Product or Plant Effect Low 產品或 工廠 影響 低 | 2-3 | Very high | 8-10 | Low-Very low | 7-10 | M
|  |  |  | Moderate | 5-6 | M
|  |  |  | High | 2-4 | L
|  |  |  | Very high | 1 | L
|  | High | 6-7 | Low-Very low | 7-10 | L
|  |  |  | Moderate | 5-6 | L
|  |  |  | High | 2-4 | L
|  |  |  | Very high | 1 | L
|  | Moderate | 4-5 | Low-Very low | 7-10 | L
|  |  |  | Moderate | 5-6 | L
|  |  |  | High | 2-4 | L
|  |  |  | Very high | 1 | L
|  | Low | 2-3 | Low-Very low | 7-10 | L
|  |  |  | Moderate | 5-6 | L
|  |  |  | High | 2-4 | L
|  |  |  | Very high | 1 | L
|  | Very low | 1 | Very high - Very low | 1-10 | L
No discernible Effect 無法辨別 影響 | 1 | Very low - Very high | 1-10 | Very high - Very low | 1-10 | L

## Flowchart & P-Diagram
P 參數圖

## PFMEA
Process Failure Mode and Effects Analysis (PFMEA)
Planning and Preparation (STEP 1 計劃與準備)
| Company Name: | ABC 公司 | Subject: | Power Supply Assembly
| Engineering Location: | 台北市 Taiwan | PFMEA Start Date: | Nov.10 , 2023 | PFMEA ID Number: | PFMEA-1
| Customer Name: | EFG 客戶 | PFMEA Revision Date: | Nov.30 , 2024 | Design Responsibility: | PE
| Model Year¦ Platform: | Medical Adapter 90W | Cross-Functional Team: | RD, QC, PE | Confidentiality Level: | Confidential
| STRUCTURE ANALYSIS (STEP 2 結構分析) |  |  | FUNCTION ANALYSIS (STEP 3 功能分析) |  |  | FAILURE ANALYSIS (STEP 4 失效分析) |  |  |  | RISK ANALYSIS (STEP 5 風險分析) |  |  |  |  | OPTIMIZATION (STEP 6 優化) | [12 empty] | CONTINUOUS IMPROVEMENT 持續改善
Issue # 議題號碼 | 1. Process Item System, Subsystem, Part Element or Name of Process 上一階系統 | 2. Process Step Station No. and Name of Focus Element 標的製程步驟 站別 | 3. Process Work Element 4M Type 下一階工作要素 4M: 人, 機, 料, 法, 環 | 1. Function of the Process Item Function of System, Subsystem, Part Element or Process 上一階系統功能 | 2. Function of the Process Step and Product Characteristic (Quantitative value is optional) 標的製程步驟 功能與產品特性 (可用計量值特性) | 3. Next Lower Level Function and Requirement of Characteristic 工作要素 4M 之功能與需求特性 | 1. Failure Effects (FE) to the Next Higher Level Element and/or End User 上一階失效效應 | 2. Failure Mode (FM) of the Focus Element 標的製程步驟失效模式 | 3. Failure Cause (FC) of the work Element 失效之原因(4M) | Original Severity (S) of FE 嚴重度 | Current Prevention Control (PC) of FC 目前預防管制 | Occurrence (O) of FC 發生度 | Current Detection Control (DC) of FC or FM 目前偵測管制 | Detection (D) of FC/FM 難檢度 | PFMEA AP 優先行動 | PFMEA Prevention Action 未來預防行動 | PFMEA Detection Action 未來偵測行動 | Responsible Person 負責單位 | Target Completion Date 預計完成日 | Status: (Open, Completed, Discarded) 狀態 | Action Taken with Pointer to Evidence 佐證 | Completion Date 實際完成日 | Severity (S) | Occurrence (O) | Detection (D) | PFMEA AP | 特殊特性 Special Characteristics | Remarks 備註 | Issue # | History/Change Authorization 歷史資料/授權
1 | 電子馬達組裝線 | [OP30]軸承鴨裝製程 | 操作人員 | 本流程將機械軸裝入外殼, 將來用以裝入車門, 得以升起與下降車窗 | 壓裝燒結軸承, 實現每次壓裝在電機殼內之位置空間, 有 3.0 +-0.5 mm | 安裝固定入卡槽定位 | 間隙太小, 造成尺寸無法使用 | 燒結輪軸向定位無法定位 | 操作人員未安裝置卡榫 | 9 |  | 2 | IPQC 抽檢 | 6 | M | 防呆治具 | IQC 抽檢 | RD | 2024/9/31 | Open |  |  | 9 | 1 | 5 | L
2 | 電子馬達組裝線 | [OP30]軸承鴨裝製程 | 壓裝機 | 本流程將機械軸裝入外殼, 將來用以裝入車門, 得以升起與下降車窗 | 壓裝燒結軸承, 實現每次壓裝在電機殼內之位置空間, 有 3.0 +-0.5 mm | 溫度上升至150 +- 1 度熱壓 | 間隙太大, 造成尺寸無法使用 | 燒結輪軸向定位無法定位 | 機台加熱塊失效 | 9 |  | 2 | PM 人員預防保養 | 6 | M | 機台自動迴路 | 機台自動監控 | PE | 2024/9/31 | Open |  |  | 9 | 1 | 3 | L
3 | 電子馬達組裝線 | [OP30]軸承鴨裝製程 | 螺帽 | 本流程將機械軸裝入外殼, 將來用以裝入車門, 得以升起與下降車窗 | 壓裝燒結軸承, 實現每次壓裝在電機殼內之位置空間, 有 3.0 +-0.5 mm | 螺帽毛邊 | 間隙太小, 造成尺寸無法使用 | 燒結輪軸向定位無法定位 | 原料毛邊 | 9 |  | 4 | IQC 抽檢 | 6 | H | 供應商換模 Checklist | IQC 機台檢測機 | IQC | 2024/9/31 | Open |  |  | 9 | 2 | 4 | L
4 | 機殼組裝 | [OP40]超音波焊接 | 操作人員 | 本流程將機殼鎖固, 防水與防防塵 | 提供外部結構強度保護、符合規格之尺寸與外觀。 | 正確設定參數 | 漏水 | 外殼破裂造成內部受外力損毀，造成功能異常 | 參數設定錯誤 | 8 | 領班 Checklist | 3 | IPQC 抽檢 | 6 | M | 機台 MES 自動設定 | Worse case 之測試評估/依設計需求落實機殼零件厚度與材質之進料檢驗 | PE | 2024-11-02 | Open |  |  | 8 | 3 | 2 | L |  |  |  | CIT Program
5 | 機殼組裝 | [OP40]超音波焊接 | 超音波機 | 本流程將機殼鎖固, 防水與防防塵 | 提供外部結構強度保護、符合規格之尺寸與外觀。 | 穩定電流 | 漏水 | 外殼破裂造成內部受外力損毀，造成功能異常 | 電流不穩 | 8 | 領班 Checklist | 2 | PM 人員預防保養 | 6 | M | 機台自動監控 | 熔接流程、外觀與尺寸檢驗標準化 | PE | 2024-11-02 | Open |  |  | 8 | 2 | 2 | L | *

## PFMEA Severity
Product General Evaluation Criteria Severity S 產品通用評估標準 嚴重性 S
Potential Failure Effects rated according to the criteria below 根據以下標準對潛在故障影響進行評級 |  |  |  |  | Blank until filled in by user 留空，直到使用者填寫為止
S | Effect 影響 | Impact to your Plant 對本廠影響 | Impact to Ship-to Plant 對下游組裝影響 | Impact to End User 對最終用戶的影響 | Corporate or Product Line Examples
10 | High 高 | Failure may result in an acute health and/or safety risk for the manufacturing or assembly worker 否則可能會對製造或裝配工人造成嚴重的健康和/或安全風險 | Failure may result in an acute health and/or safety risk for the manufacturing or assembly worker 否則可能會對製造或裝配工人造成嚴重的健康和/或安全風險 | Affects safe operation of the vehicle and/or other vehicles, the health of driver or passenger(s) or road users or pedestrians. 影響車輛和/或其他車輛的安全操作、操作人員或乘客或道路消費者或行人的健康。
9 |  | Failure may result in in-plant regulatory noncompliance 否則可能導致廠內不符法規 | Failure may result in in-plant regulatory noncompliance 否則可能導致廠內不符法規 | Noncompliance with regulations. 不符合法規。
8 | Moderately high 中高 | 100% of production run affected may have to be scrapped. Failure may result in in-plant regulatory noncompliance or may have a chronic health and/or safety risk for the manufacturing or assembly worker 受影響的 100% 生產運行可能不得不報廢。如果不這樣做，可能會導致廠內不遵守法規，或者可能對製造或裝配工人造成長期的健康和/或安全風險 | Line shutdown greater than full production shift; stop shipment possible; field repair or replacement required (Assembly to End User) other than for regulatory noncompliance. Failure may result in in-plant regulatory noncompliance or may have a chronic health and/or safety risk for the manufacturing or assembly worker. 生產線停工超過整個生產班次;可能停止裝運;需要現場維修或更換（組裝到最終使用者），但不符合法規規定的情況除外。否則可能導致廠內不符法規，或者可能對製造或裝配工人造成長期的健康和/或安全風險。 | Loss of primary vehicle function necessary for normal driving during expected service life. 在預期的使用壽命內，正常駕駛所必需的主要車輛功能喪失。
7 |  | Product may have to be sorted and a portion (less than 100%) scrapped; deviation from primary process; decreased line speed or added manpower 產品可能需要進行分揀，並且一部分（小於 100%）報廢;主流程的偏差;降低生產線速度或增加人力 | Line shutdown from 1 hour up to full production shift; stop shipment possible; field repair or replacement required (Assembly to End User) other than for regulatory noncompliance 生產線停機 1 小時直至完全生產班次;可能停止裝運;需要現場維修或更換（組裝到最終使用者），但不符合法規的情況除外 | Degradation of primary vehicle function necessary for normal driving during expected service life. 在預期的使用壽命內，正常駕駛所必需的主要車輛功能退化。
6 | Moderately low 中低 | 100% of production run may have to be reworked off line and accepted 100% 生產運行可能必須要求下線且允收 | Line shutdown up to one hour 當線停機長達一小時 | Loss of secondary vehicle function. 輔助車輛功能喪失
5 |  | A portion of the production run may have to be reworked off line and accepted 部分生產運行可能必須要求下線且允收 | Less than 100% of product affected; strong possibility for additional defective product sort required; no line shutdown 少於 100% 的產品受到影響;需要對有缺陷的產品進行額外的重工可能性很大;不停線 | Degradation of secondary vehicle function. 輔助車輛功能退化
4 |  | 100% of production run may have to be reworked in station before it is processed 100% 的生產運行可能必須在工作站進行返工，然後才能進行加工 | Defective product triggers significant reaction plan; additional defective products not likely; sort not required 有缺陷的產品觸發重大反應計劃;不太可能有其他缺陷的產品;重工不需要 | Very objectionable appearance, sound, vibration, harshness, or haptics. 非常令人反感的外觀、聲音、振動、粗糙度或觸覺。
3 | Low 低 | A portion of the production run may have to reworked in-station before it is processed. 在進行處理之前，生產運行的一部分可能必須在站內進行返工。 | Defective product triggers minor reaction plan; additional defective products not likely; sort not required 有缺陷的產品但不需觸發反應計劃;不太可能有其他缺陷;清倉不需要 | Moderately objectionable appearance, sound, vibration, harshness, or haptics. 中度令人反感的外觀、聲音、振動、粗糙度或觸覺。
2 |  | Slight inconvenience to process, operation, or operator 對過程、操作或操作員的輕微不便 | Defective product triggers no reaction plan; additional defective products not likely; sort no required; requires feedback to supplier 有缺陷的產品但不需觸發反應計劃;不太可能有其他缺陷;清倉不需要, 但須向供應商反饋 | Slightly objectionable appearance, sound, vibration, harshness, or haptics 輕微令人反感的外觀、聲音、振動、粗糙度或觸覺
1 | Very low 極低 | No discernible effect 沒有明顯的影響 | No discernible effect or no effect 沒有明顯的影響或無影響 | No discernible effect 沒有明顯的影響

## PFMEA Occurrence
Occurrence Potential (O) for the Process 流程的發生可能性 (O)
Potential Failure Causes rated according to the criteria below. Consider Prevention Controls when determining the best Occurrence estimate. Occurrence is a predicted qualitative rating made at the time of evaluation and may not reflect the actual occurrence, The occurrence rating number is a relative rating within the scope of the FMEA (process being evaluated). For Prevention Controls with multiple Occurrence Ratings, use the rating that best reflects the robustness of the control. 潛在故障原因根據以下標準進行評級。在確定最佳發生估計值時，請考慮預防控制。發生率是在評估時做出的預測定性評級，可能與實際發生情況不符，發生率評級數位是 FMEA（被評估流程）範圍內的相對評級。對於具有多個發生評級的預防控制，請使用最能反映控制穩健性評級的評級。 |  |  |  | Blank until filled in by user 留空，直到使用者填寫為止
O | Prediction of Failure Cause Occurring | Type of Control | Prevention Controls 預防控制 | Corporate or Product Line Examples
10 | Extremely high 極高 | None 無 | No prevention controls 無預防管制
9 | Very high 非常高 | Behavioral 行為 | Prevention controls will have little effect in preventing failure cause. 預防控制在防止故障原因方面效果不大。
8
7 | High 高 | Behavioral or Technical 行為或技術 | Prevention controls somewhat effective in preventing failure cause. 預防控制在一定程度上有效地防止了故障原因。
6
5 | Moderate 中 |  | Prevention controls are effective in preventing failure cause. 預防控制可有效防止故障原因。
4
3 | Low 低 | Best Practices; Behavioral or Technical 最佳實踐;行為或技術 | Prevention controls are highly effective in preventing failure cause. 預防控制在防止故障原因方面非常有效。
2 | Very Low 非常低
1 | Extremely Low 極低 | Technical 技術 | Prevention controls are extremely effective in preventing failure cause from occurring due to design (i.e. part geometry) or process (i.e. fixture or tooling design). Intent of prevention controls - Failure Mode cannot be physically produced due to the Failure Cause. 預防控制在防止由於設計（即零件幾何形狀）或工程（即夾具或工具設計）而導致的故障原因方面非常有效。預防控制的意圖 - 由於故障原因，無法物理生成故障模式。

## PFMEA Detection
Detection Potential (D) for the Validation of the Process Design 流程通用評估標準 難檢性 D
Detection Controls rated according to the Detection Method Maturity and Opportunity for Detection. 偵測控制根據偵測方法成熟度和偵測機會進行評級。 |  |  |  | Blank until filled in by user 留空，直到使用者填寫為止
D | Ability to Detect 偵測能力 | Detection Method Maturity 偵測方法成熟度 | Opportunity for Detection 偵測機會 | Corporate or Product Line Example
10 | Very Low 非常低 | No testing or inspection method has been established or is known. 尚未建立或已知的測試或檢查方法。 | The failure mode will not or cannot be detected. 不會或無法檢測到故障模式。
9 |  | It is unlikely that the testing or inspection method will detect the failure mode. 測試或檢查方法不太可能檢測到故障模式。 | The failure mode is not easily detected through random or sporadic audits. 通過隨機或零星的稽核不容易檢測到故障模式。
8 | Low 低 | Test or inspection method has not been proven to be effective and reliable (i.e. plant has little or no experience with method, gauge R&R results marginal on comparable process or this application, etc.). 測試或檢查方法尚未被證明是有效和可靠的（即工廠在方法、量規 R&R 結果在流程或此應用等方面的經驗很少或根本沒有經驗）。 | Human inspection (visual, tactile, audible), or use of manual gauging (attribute or variable) that should detect the failure mode or failure cause. 人工檢查（視覺、觸覺、聽覺）或使用手動測量（屬性或變數），應該大致可檢測出故障模式或故障原因。
7 |  |  | Machine-based detection (automated or semi-automated with notification by light, buzzer, etc.), or use of inspection equipment such as a coordinate measuring machine that should detect failure mode or failure cause. 基於機器的檢測（自動或半自動，通過燈、蜂鳴器等發出通知），或使用檢查設備，例如量測機應大致可測出故障模式或故障原因。
6 | Moderate 中 | Test or inspection method has been proven to be effective and reliable (i.e. plant has experience with method, gauge R&R results are acceptable on comparable process or this application, etc.). 測試或檢查方法已被證明是有效和可靠的（即工廠具有方法經驗，治具R&R結果在此流程或此應用等上是可以接受的）。 | Human inspection (visual, tactile, audible), or use of manual gauging (attribute or variable) that will detect the failure mode or failure cause (including product sample checks). 人工檢查（視覺、觸覺、聽覺），或使用手動測量（屬性或變數）來檢測故障模式或故障原因（包括產品樣品檢查）。
5 |  |  | Machine-based detection (semi-automated with notification by light, buzzer, etc.), or use of inspection equipment such as a coordinated measuring machine that will detect failure mode or failure cause (including product sample checks). 基於機器的檢測（半自動，燈、蜂鳴器等進行通知），或使用檢查設備，例如檢測故障模式或故障原因的測量機（包括產品樣品檢查）。
4 | High 高 | System ha been proven to be effective and reliable (i.e. plant has experience with method on identical process or this application), gauge R&R results are acceptable, etc. 系統已被證明是有效和可靠的（即工廠具有在相同過程或此應用上的方法經驗），治具R&R結果是可以接受的。 | Machine-based automated detection method that will detect the failure mode downstream, prevent further processing or system will identify the product as discrepant and allow it to automatically move forward in the process until the designated reject unload area. Discrepant product will be controlled by a robust system that will prevent outflow of the product from the facility. 基於機器的自動檢測方法，在下游站找出故障模式，防止進一步流出，或者系統能識別差異產品，並允許其在流程中自動向前檢出，排除到指定的不合格品卸料區域。不良產品能被一個強大的系統控制，該系統將防止不良產品從設備中流出。
3 |  |  | Machine-based automated detection method that will detect the failure mode in-station, prevent further processing or system will identify the product as discrepant and allow it to automatically move forward in the process until the designated reject unload area. Discrepant product will be controlled by a robust system that will prevent outflow of the product from the facility. 基於機器的自動檢測方法，能檢測當站內的故障模式，防止進一步處理，或者系統能識別產品不良，並允許其在流程中自動向前檢出，排除到指定的不合格品卸料區域。不良產品能被一個強大的系統控制，該系統將防止不良產品從設備中流出。
2 |  | Detection method has been proven to be effective and reliable (i.e. plant has experience with method, error-proofing verifications, etc.) 檢測方法已被證明是有效和可靠的（即工廠具有方法、防錯驗證等方面的經驗） | Machine-based detection method that will detect the cause and prevent the failure mode (discrepant part) from being produced. 基於機器的檢測方法，可檢測原因並防止產生故障模式（不符合之料件）。
1 | Very high 非常高 | Failure mode cannot be physically produced as-designed or processed, or detection methods proven to always detect the failure mode or failure cause. 失效模式不會按設計或處理實際產生，偵測方法有效能偵測失效模式或失效原因。

## PFMEA AP
Action Priority (AP) for DFMEA and PFMEA
Action Priority is based on combinations of Severity, Occurrence, and Detection ratings in order to prioritize actions for risk reduction. 行動優先順序是基於嚴重度S、發生度O和難檢度D評級的組合 為了確定降低風險的優先行動的優先次序。 | [6 empty] | Blank until filled in by user 留空，直到使用者填寫為止
Effect 效應 | S | Prediction of Failure Cause Occurring 預測 失敗原因 發生 | O | Ability to Detect 偵測能力 | D | ACTION PRIORITY (AP) 行動優先 | Comments
Product or Plant Effect Very high 產品或 工廠 影響 非常高 | 9-10 | Very high | 8-10 | Low-Very low | 7-10 | H
|  |  |  | Moderate | 5-6 | H
|  |  |  | High | 2-4 | H
|  |  |  | Very high | 1 | H
|  | High | 6-7 | Low-Very low | 7-10 | H
|  |  |  | Moderate | 5-6 | H
|  |  |  | High | 2-4 | H
|  |  |  | Very high | 1 | H
|  | Moderate | 4-5 | Low-Very low | 7-10 | H
|  |  |  | Moderate | 5-6 | H
|  |  |  | High | 2-4 | H
|  |  |  | Very high | 1 | M
|  | Low | 2-3 | Low-Very low | 7-10 | H
|  |  |  | Moderate | 5-6 | M
|  |  |  | High | 2-4 | L
|  |  |  | Very high | 1 | L
|  | Very low | 1 | Very high - Very low | 1-10 | L
Product or Plant Effect High 產品或 工廠 影響 高 | 7-8 | Very high | 8-10 | Low-Very low | 7-10 | H
|  |  |  | Moderate | 5-6 | H
|  |  |  | High | 2-4 | H
|  |  |  | Very high | 1 | H
|  | High | 6-7 | Low-Very low | 7-10 | H
|  |  |  | Moderate | 5-6 | H
|  |  |  | High | 2-4 | H
|  |  |  | Very high | 1 | M
|  | Moderate | 4-5 | Low-Very low | 7-10 | H
|  |  |  | Moderate | 5-6 | M
|  |  |  | High | 2-4 | M
|  |  |  | Very high | 1 | M
|  | Low | 2-3 | Low-Very low | 7-10 | M
|  |  |  | Moderate | 5-6 | M
|  |  |  | High | 2-4 | L
|  |  |  | Very high | 1 | L
|  | Very low | 1 | Very high - Very low | 1-10 | L
Product or Plant Effect Moderate 產品或 工廠 影響 中等 | 4-6 | Very high | 8-10 | Low-Very low | 7-10 | H
|  |  |  | Moderate | 5-6 | H
|  |  |  | High | 2-4 | M
|  |  |  | Very high | 1 | M
|  | High | 6-7 | Low-Very low | 7-10 | M
|  |  |  | Moderate | 5-6 | M
|  |  |  | High | 2-4 | M
|  |  |  | Very high | 1 | L
|  | Moderate | 4-5 | Low-Very low | 7-10 | M
|  |  |  | Moderate | 5-6 | L
|  |  |  | High | 2-4 | L
|  |  |  | Very high | 1 | L
|  | Low | 2-3 | Low-Very low | 7-10 | L
|  |  |  | Moderate | 5-6 | L
|  |  |  | High | 2-4 | L
|  |  |  | Very high | 1 | L
|  | Very low | 1 | Very high - Very low | 1-10 | L
Product or Plant Effect Low 產品或 工廠 影響 低 | 2-3 | Very high | 8-10 | Low-Very low | 7-10 | M
|  |  |  | Moderate | 5-6 | M
|  |  |  | High | 2-4 | L
|  |  |  | Very high | 1 | L
|  | High | 6-7 | Low-Very low | 7-10 | L
|  |  |  | Moderate | 5-6 | L
|  |  |  | High | 2-4 | L
|  |  |  | Very high | 1 | L
|  | Moderate | 4-5 | Low-Very low | 7-10 | L
|  |  |  | Moderate | 5-6 | L
|  |  |  | High | 2-4 | L
|  |  |  | Very high | 1 | L
|  | Low | 2-3 | Low-Very low | 7-10 | L
|  |  |  | Moderate | 5-6 | L
|  |  |  | High | 2-4 | L
|  |  |  | Very high | 1 | L
|  | Very low | 1 | Very high - Very low | 1-10 | L
No discernible Effect 無法辨別 影響 | 1 | Very low - Very high | 1-10 | Very high - Very low | 1-10 | L